## Slide 1
HIM 2009 spring
Development of the Scintillation-fiber Detector
for the Neutron-Beam Profile Measurement
C. Kim,* B. Hong, R. J. Hu, M. Jo, K. S. Lee, S. Park, and K. S. Sim
Department of Physics and Korea Detector Laboratory,
Korea University, Seoul 136-701

## Slide 2
Contents
Motivation
Detector Characteristics
  Scintillation fiber
  Si-photodiode array and Electronics
  Protection system
  Control device for scan operations
Test
  Calibration by X-ray
  Neutron-beam profile measurement at KIRAMS
Beam-profile images in absorbed dose rates
Detector modification
Conclusions and Prospects

## Slide 3
1. Motivation
High LET (linear energy transfer) radiations:
  - Fast neutrons, protons, heavy-ions…
  - significantly large energy deposit in tissue per unit length
Fast-neutron radiotherapy:
  Very effective treatment for soft-tissue cancer
Necessity of precise neutron beam-profile
measurement device
We designed, built, and tested scintillation fiber detector based on a current-mode
electronics.
1.
2.
a.
b.
c.
d.
3.
a.
b.
4.
5.
6.
1/19

## Slide 4
2. Detector characteristics - a. Scintillation fiber
11.8 mm
59.7 mm
9.0 mm
46.2 mm × 3.0 mm
(46 × 3 pieces)
Single-clad scintillation fiber (Bicron BCF-60)
9 mm fiber length:
  - Compromised between neutron sensitivity enhancement &
  multiple scattering probability suppression
Three-layer structure for maximized light yield produced
by the scattered protons
9.0 mm
1.0 × 1.0 mm2
Neutron
Emitted photon
(Green, 530 nm)
Energetic
proton
Neutron
Energetic
proton
Line scan detector with 46 × 3 pixels
1.
2.
a.
b.
c.
d.
3.
a.
b.
4.
5.
6.
2/19

## Slide 5
Electronics board
Scintillation fibers
46 channel Si-photodiode array
(Hamamatsu S4111-46Q)
Electronics board
46 channel Si-photodiode array
(Hamamatsu S4111-46Q)
46 channel Si-photodiode (Hamamatsu S4111-46Q)
  - Dark current: ~ 10 pA
Current-integration-mode electronics
  - Designed for high-intensity fast neutron-beams
  (neutron beams for typical radiotherapies, 108 ~ 1010 Hz/cm2 )
3/14
2. Detector characteristics - b. Si-photodiode array and Electronics
1.
2.
a.
b.
c.
d.
3.
a.
b.
4.
5.
6.
3/19

## Slide 6
2. Detector characteristics - b. Si-photodiode array and Electronics
1.
2.
a.
b.
c.
d.
3.
a.
b.
4.
5.
6.
4/19

## Slide 7
LV bias
(+5V)
Silicon
photo -
diode
Preamp
ADC
Digital
processor
Data
buffer
MCU
Ethernet
link
Data
control
in PC
Output pulses fed into preamp (preamp sensitivity: 10-10 ~ 10-2 A)
Preamp converts output pulses to voltage-sensitive pulses
Preamp outputs are multiplexed to four ADCs
ADCs convert the voltage-sensitive analog inputs into 12-bit digital signals (in every 7.8 μs) and feed them to a digital processor
Input signals transferred from the ADCs are summed over 1 ms and fed into a data buffer area
Data from digital processor are transferred to PC by MCU via Ethernet link
2. Detector characteristics - b. Si-photodiode array and Electronics
1.
2.
a.
b.
c.
d.
3.
a.
b.
4.
5.
6.
5/19

## Slide 8
  2. Detector characteristics - c. Protection system
V shape collimator
(Gd2O3 + epoxy mixture)
Acryl plate
(1 cm thickness)
Gd2O3 powder
(1 mm thickness)
Upper collimator
(Gd2O3 + epoxy mixture)
Glass plate
(2 mm thickness)
Two-step collimation for the fast neutrons
  - Collimators made of natural Gd2O3 powder and epoxy
  - Slit width of V-shape collimator: 1 mm
Gd layers for fast/thermal neutron shielding
1.
2.
a.
b.
c.
d.
3.
a.
b.
4.
5.
6.
6/19

## Slide 9
  2. Detector characteristics - d. Control device for scan operations
Transverse movement (red line):
  - continuous with adjustable speed (1~5 cm/s)
Vertical movement (yellow line): 46 mm per step
Total possible scan area: 30 × 32 cm2
1.
2.
a.
b.
c.
d.
3.
a.
b.
4.
5.
6.
7/19

## Slide 10
Calibration test for the sensitivity of each channel
Test was performed by the X-ray generator (8 mA, 70 kV)
  - The portable X-ray gun was placed at three different distances
  (8, 10, and 16 cm) from the detector
Results:
- Sensitivity varies at the level of a few %
- Channel responses were measured precisely.
- Calibration factors were obtained for the channel responses.
  3. Test - a. Calibration by X-ray
1.
2.
a.
b.
c.
d.
3.
a.
b.
4.
5.
6.
8/19

## Slide 11
  3. Test - b. Neutron-beam profile measurement at KIRAMS
Fast-neutron beam was provided by MC50 cyclotron at KIRAMS
Distance from Be target to the scan detector: 123 cm
Incident proton energy: 45 MeV
Be target thickness: 10.5 mm
Test was performed with two beam currents: 10 μA and 20 μA
9/19
1.
2.
a.
b.
c.
d.
3.
a.
b.
4.
5.
6.

## Slide 12
  3. Test - b. Neutron-beam profile measurement at KIRAMS
20 μA, 123 cm, 8×8 cm2, Aug25, PM0315
Channels
time (s)
Q (pC)
Q (pC)
1.
2.
a.
b.
c.
d.
3.
a.
b.
4.
5.
6.
10/19

## Slide 13
  3. Test - b. Neutron-beam profile measurement at KIRAMS
11/19
Channels
time (s)
Q (pC)
Q (pC)
10 μA, 123 cm, 8×8 cm2, Aug25, PM0328
1.
2.
a.
b.
c.
d.
3.
a.
b.
4.
5.
6.

## Slide 14
  3. Test - b. Neutron-beam profile measurement at KIRAMS
Beam-profile measurement conditions:
- Corresponding mean charge value to neutron beam intensities:
~ 4.0 pC ↔ ~ 4.8 × 107 Hz/cm2 (20 μA beam current)
~ 1.8 pC ↔ ~ 2.4 × 107 Hz/cm2 (10 μA beam current)
- Test performed under the beam intensities lower than operational intensity
(108 ~ 1010 Hz/cm2)
Measured beam profiles at KIRAMS (left) and Unfolded neutron flounce spectrum
(lethargy spectrum) measured by Boner sphere system (KAERI/RR-2442/2003) (right)
1.
2.
a.
b.
c.
d.
3.
a.
b.
4.
5.
6.
12/19

## Slide 15
  4. Beam-profile images in absorbed dose rates
13/19
Where,
Ap : area of each image pixel (cm2)
ip : proton beam current (nA)
Ei : neutron energy (MeV)
Ei ΔФi /ΔEi : differential neutron flux (n/s/nA)
δEi : energy deposited in the detector (GEANT4)
єi : interaction sensitivity (GEANT4)
Qm : mean charge induced in the beam area
1) Beam-profile image (charge distribution)
- Detector signal is roughly proportional to the deposited energy in the detector
- Beam-profile images ≈ distribution of expose dose rate induced in the
detector
2) Conversion by using Geant4 simulations
→ The conversion factor
∴ Distribution of the beam profile in Gy :
Neutron spectrum measured at 1.5 m (n/s/nA) (KAERI/RR-2442/2003)
1.
2.
a.
b.
c.
d.
3.
a.
b.
4.
5.
6.

## Slide 16
3) Applying to human-body:
- Proton beam status: 45 MeV, 20 μA
- Be target thickness: 10.5 mm
- Distance: 123 cm
- Beam area : 8 × 8 cm2
- Human body phantom : 30 × 30 × 20 (depth) cm3
- Size of a voxel : 1 cm3
  4. Beam-profile images in absorbed dose rates
1.
2.
a.
b.
c.
d.
3.
a.
b.
4.
5.
6.
14/19

## Slide 17
  4. Beam-profile images in absorbed dose rates
1.
2.
a.
b.
c.
d.
3.
a.
b.
4.
5.
6.
15/19

## Slide 18
~70 mm
1.0 × 1.0 mm2
Emitted photon
(Green, 530 nm)
Energetic
proton
Neutron
Bended scintillation fibers
(46 × 2 pieces)
47.0 mm × 2.0 mm
55 mm
41 mm
60.0 mm
15.0 mm
10.0 mm
10.0 mm
10.0 mm
  5. Detector modification
Bended structure to avoid direct beam-exposure
Improvement in statistics: {(108 → 107) ~ 1010} (Hz/cm2)
- Increased fiber length (9 mm → ~ 20 mm)
- Collimator slit removed (open width: 1 mm → 2 mm)
- * movement speed will also be adjusted
1.
2.
a.
b.
c.
d.
3.
a.
b.
4.
5.
6.
16/19

## Slide 19
mounted Electronics
2 mm thickness
Gd + Epoxy mixture
5 mm thickness
acryl plate
10 mm thickness
acryl plate
~ 15 mm thickness
Gd + Epoxy mixture
30~40 mm thickness
acryl plate
  5. Detector modification
1.
2.
a.
b.
c.
d.
3.
a.
b.
4.
5.
6.
17/19

## Slide 20
  6. Conclusions and Prospects
Conclusions
Fast-neutron beam profile was successfully measured.
The detector composed of scintillation fiber and current-integration mode electronics was proved as a reliable
instrument for the beam-profile measurement.
High-precision data measurement:
not only beam area but also beam-halo were observed.
Stable beam-profile measurement is possible
Direct result (charge distribution) can be converted to absorbed dose rates by Geant4 simulation
18/19
1.
2.
a.
b.
c.
d.
3.
a.
b.
4.
5.
6.

## Slide 21
19/19
  6. Conclusions and Prospects
Prospects
Second neutron-beam test arranged at KIRAMS: March 16
Further, and deeper simulations by Geant4 are required for precise absorbed/equivalent dose rates calculation
1.
2.
a.
b.
c.
d.
3.
a.
b.
4.
5.
6.

## Slide 22
Thank you!

## Slide 23
  Backup Slides - Signal produce
Elastic scatterings
  n-p elastic scattering: deposited energy by recoiled proton
  (dominant)
  n-12C elastic scattering: deposited energy by recoiled 12C
  Neutron signal is proportional to deposited energy
  The beam-profile images directly reflects the distribution of
  absorbed doses
Inelastic scatterings
  p(n , γ)d → signals by secondary deuteron
  n + 12C → 3 α’s

## Slide 24
  Backup Slides - Gadolinium
동위원소 질량비 |  | 내부전환에 의해 발생하는 전자의 에너지 | 
질량수 | 존재비 (%) | Energy (KeV) | 내부전환 확률 (%)
152 | 0.2 | 29 | 9.82
154 | 2.2 | 39 | 4.19
155 | 14.8 | 71 | 26.80
156 | 20.5 | 78 | 6.17
157 | 15.7 | 81 | 4.97
158 | 24.8 | 88 | 1.16
160 | 21.8 | 131 | 3.41
여백 |  | 150 이상 | ~2
157Gd: 6.3 × 106 barn
155Gd: 3.6 × 105 barn

## Slide 25
  Backup Slides – Detector arm
Scan Device
Beam Area
Detector arm
Detector

## Slide 26
  Backup Slides – (tagging efficiency × deposited energy) / neutron

## Slide 27
  Backup Slides – Counts of secondary particles / neutron

## Slide 28
  Backup Slides - Position sensitivity test
PMT
Pb Brick
Pb Brick
50 mm
60 mm
γ rays
Co60 (~1.25 MeV)
~10 mm
pixel detector (side view)
~ 46 mm
Schematic diagram of position sensitivity test. The width of γ-ray pass (open space
between lead bricks) adjusted to overcome low irradiation rate of γ-source.
4 points (by 1 cm interval) of detector
checked to examine position-sensitivity
of detector.
γ-ray distribution of each points were
all uniform.

## Slide 29
  Backup Slides - LED Test
Amounts of photon correspondent to LED power measured by PMT
Same experiment repeated by designed electronics
PMT
Simple circuit for LED power
Electronics
Si-photodiode array
low-brightness LED (green, ~550 nm)
2 pieces of obscure glasses
taped paper cylinder
paper cup
power supply
~7 cm
~3 cm
Schematic diagram of LED test