## Slide 1: Performance of the X-ray PSD based on the MWPC
Ryu Min Sang
Neutron Science Division
Korea Atomic Energy Research Institute (KAERI)
Position sensitive detector (PSD)
Multi-wire proportional chamber (MWPC)
2010.02.25-27 Heavy Ion Meeting
Yongpyong resort, Pyungchang

## Slide 2: Contents
Introduction
  X-ray diffraction
  Small angle X-ray scattering (SAXS)
  Position sensitive detector (PSD)
Detector construction
DAQ and test setup
Experimental results
Fabricated PSD at KAERI
Summary

## Slide 3: Introduction – X-ray
In 1895, Wilhelm Conrad Roentgen (GE) discovered the X-ray.
< general properties >
X-ray detection method
  Imaging
  Fluorescence (ZnS, CdS, NaI)
  Ionization
Light speed in vacuum
Diffraction like particle
Reflection index ~ 1
(impossible to focus the X-ray)
medical imaging & material experiment due to the high transparency

## Slide 4: Introduction – X-ray diffraction
In 1912, Laue (GE) succeeded the X-ray diffraction experiment .
X-ray diffracts when its wavelength is same as the distance in between atoms.
In 1913, Bragg (UK) makes Laue’s diffraction formula more simply.
Constructive interference: n = integer
d: distance between the interface of atoms
λ : wavelength
θ : incidence angle

## Slide 5: Introduction - SAXS
Small angle X-ray scattering (SAXS):
  - research the shape and size of the high molecular materials.
- role of PSD: to measure the scattered pattern.
q: scattering vector
d: distance between the surface of atoms
λ : wavelength
2θ : scattering angle
http://mrl.uscb.edu/~safinyaweb/XRD.htm

## Slide 6: SAXS at Hannam University
Energy range of photon for diffraction and scattering experiment: 5 ~ 30 keV
< X-ray generator >
Rigaku rotating anode
3 kW filament
40 kV, 70 mA
< Focusing mirror >
Osmic
Confocal Max-Flux
Source-focus distance: 1200 mm
Photo-electric effect
Wavelength of X-ray: 1.54 Å at 8.04 keV
2D MWPC
Vacuum pipe
sample
slit2
slit1
Focusing mirror
X-ray generator
MWPC FEE

## Slide 7: Introduction - PSD
Multi-wire proportional chamber (MWPC)
- single counting method
<strength>
- high X-ray detection efficiency
- good position linearity
- geometrical flexibility
- fast DAQ
- low production cost
<weakness>
- relatively poor position resolution compared to SDAD.
- poor rate capability
Silicon diode array detector (SDAD) or charge-coupled device (CCD)
- charge accumulation method
<strength>
- good rate capability
- good position resolution
<weakness>
- high background
- narrow dynamic range
- poor time resolution
- slow DAQ
Why is MWPC attracted than the semiconductor detector?

## Slide 8: Operation of X-ray PSD
Energy range of photon for diffraction and scattering experiment: 5 ~ 30 keV
Photo-electrical effect creates the photo-electron or Auger electron.
Average ionization energy of noble gas (Ar, Kr, Xe): ~ 30 eV
- number of produced electron-ion pairs: ~160 at 5 keV
~267 at 8 keV
~1000 at 30 keV
<MWPC>
G. Charpak et al., Nucl. Instr. and Meth. 62 262 (1968)
G. Charpak and F. Sauli, Nucl. Instr. and Meth. 162 405 (1979)
Incident X-ray
Photo-electric effect
Electron avalanche occurs toward the anode.
Collect the electrons on the anode.
Positive ion drifts toward the cathode.
Induced signal on the cathode.

## Slide 9
Effective area : 120 x 120 mm
230 mm
230 mm
0.4 mm aluminized carbon fiber
~80 % transmittance at 8 keV (1.54 Å)
Ø150 mm
Anode: Ø10 ㎛, 1.25 mm spacing.
(Ø30 ㎛ used for 1st and 2nd wires at both end side.)
Dual Cathode: Ø30 ㎛, 1.25 mm spacing
Drift plate
Anode-Cathode gap: 3.2 mm
Drift-Cathode gap: 3.2 mm
Detector configuration
Cathode (Y1, Y2)
Cathode (X1, X2)
Anode
Drift
Aluminized carbon fiber

## Slide 10: Winding Machine
Weight (wire tension control)
M3
M2
M1
Linear translator
(wire step control)
Wire
Wire frame (PCB)
*M means a stepping motor
M1 : Wire step control with linear translator
M2 : Wire frame rotator
M3 : Wire length control
Fixed pulley
Movable pulley
CERN report 77-09(1977)
M3
M2
Motor controller
M2
M3
M1
Weight (wire tension control)
M1

## Slide 11: Picture of X-ray PSD
pursing

## Slide 12: Data acquisition
Anode
X1
Detector delay time
X2
Y1
Y2
TDC delay
TDC delay
ΔX
ΔY
TDC gate time
Delay-line readout method
Inductor (145 nH) & capacitor (56 pF)
Total delay time: 170 ns
X1
X2
Y1
Y2
Constant
Fraction
Discriminator
CABLE
DELAY
ATDC
Anode
X
-
Left
X
-
Right
Position
Sensitive
Detector
IN
OUT
IN
OUT
IN
OUT
COM
USB
PA
PA
PA
EXT
count
EXT
triger
OUT
IN
OUT
IN
Y
-
Left
Y
-
Right
PA
PA
Charge
Avalanche
Anode wire
cathode wire
or
electrode
delay
delay

## Slide 13: Operational region
Ar + C2H6 + CF4 at 1.5 atm
Drift voltage: -1200 V
Fe-55: 6.4 keV, 10 mCi
(half life time : 2.744 y)
Source to PSD: 45 cm
+2475 ~ 2575V

## Slide 14: Position resolution
Ar 1.5 atm
Drift voltage: -1200 V
High voltage: +2425 ~ 2575 V
Voltage step: +50 V
Fe-55 (KAERI)
NEW
OLD
+2525 ~ 2575V

## Slide 15: Uniformity
Ar 1.5 atm
Drift voltage: -1200 V
High voltage: +2525 V
Fe-55 (KAERI)
Source to PSD: 45 cm
Time: 21600 sec (6 hours)

## Slide 16: Operational region
Am-241(103.4 uCi): ~14 keV (KAERI)
Two plots are shown the equivalent trend.
BKG rate has decreased and has been stable.
X-ray: 8.04 keV (Hannam Univ.)
(half life time : 432.6 y)
+3100 V
Xe + C2H6 + CF4 at 1.55 atm
Silver behenate (SB): standard sample for SAXS

## Slide 17: BKG & SB at SAXS
BKG rate has decreased.
18 cps at the beginning.
14 cps after 24 hours.
11 cps after 72 hours.
11.1 cps (BKG)
Xe 1.55 atm
Drift voltage: -1200 V
HV: +3050 V
3600 sec
Silver behenate (SB)

## Slide 18: Comparison
Oct. 7 (NEW)
Xe 1.55 atm
Drift voltage: -1200 V
High voltage: +3050 V
Aug. 6 (OLD)
Drift voltage: -1200 V
High voltage: +2750 V
linear
Log
The first peaks can be different due to beam focusing and fluctuation of beam flux.
The second and third peaks correspond to the old data.

## Slide 19: X-ray fluctuation
Chart: 
Counting rate (cps)
Cooler on/off
Cooler
Counting rate (cps)
Time (1bin=14sec)
Xe 1.55 atm
Drift voltage: -1200 V
360 trial
Total time: ~ 5040 sec (84 min)
High voltage: +3050 V
sample: LDPE
Counting rate shows not only some coincidence with the on/off of the cooler but also it presents unmatched period.

## Slide 20: Introduction - XICS
X-ray imaging crystal spectrometer (XICS)
- to measure ion temperatures in the tokamak device.
  Doppler-broadening
  Movement of the center wavelength (spherically bent crystal)
Doppler-Broadening: the spectral lines are broaden caused by a distribution of velocities of atoms or molecules.
Phys. Rev. Lett. 42 304 (1979) M. Bitter et al.
Rev. Sci. Instrum. 70, 1 (1999) M. Bitter et al.
J. K. Cheon, Ph.D. thesis, Kyungpook national university, (2008)
λ0 : center wavelength of the spectrum
k : Boltzmann constant
Johann configuration
Ar is injected in the tokamak.
He-like Ar (Ar XVII, Ar16+) has only two electrons due to the high energy plasma.
Ar16+ and plasma thermally equilibrated.
Ar16+ emits the X-ray.

## Slide 21: 4 segment PSD for XICS
3.9494 Å
3.9944 Å
X-ray spectrum of He-like Ar
PSD requirements
Photon counting
Enough rate capability
- Large area
Fast DAQ
position resolution (< 0.5 mm)
- Radiation hardened (fast neutron)
Ar16+ emits the X-ray (3.9494 ~ 3.9944 Å).

## Slide 22: Fabricated PSD at KAERI: part I
2D curved X-ray PSD
1D & 2D X-ray PSD

## Slide 23: Fabricated PSD at KAERI: part II
1D Neutron PSD
Neutron Monitor
2D Large Neutron PSD for SANS

## Slide 24: Summary
Many kinds of PSD have been producing at KAERI.
  Small angle X-ray scattering (SAXS)
  X-ray imaging crystal spectrometer (XICS)
  Neutron monitor
  Neutron diffraction
  Small angle neutron scattering (SANS)
PSD
  Low production cost
  Large effective area
  Geometrical flexibility
  Fast DAQ
  Good position resolution
  Sensitivity depending on gas and energy